产品单页介绍 — F2002

文档编号：	SPI-F2002-REV02-2021

产品型号：	F2002

产品名称：	程控精密电流源

关键词：	105mA/105V/11W、强电感，可调钳位电压

高精度、高稳定度、高分辨率、程控

保修服务：	1年有限保修服务

校准服务：	1次免费校准

计量服务：	非强制计量产品、付费开具计量证书

使用标准：	GB-1002、EIA-232-E、USB 1.1 Standard、IEC 825-1

产品特点：	 输出电流		 输出范围：-105mA—+105mA DC，准双极性

 步进分辨率1A（1/105000 F.S.）

 准确度：（0.015%设置值+1A）

 线性度：优于0.005%

 温度系数：优于3ppm/C

 短期稳定度：优于10ppm/h

 噪声/纹波：1A RMS典型值（低噪声模式）

 直流输出阻抗>500M

 浮动输出，I、III象限，内部继电器换向

 钳位电压		 设置范围：0—105V DC

 设置分辨率0.1V

 准确度：（0.5%设置值+0.5V）

 输出功率		 最大输出功率：11W

 自动控制		 标准配备RS-232C接口和易用的接口特性

 标准配备内部USB—RS-232C转换接口

 同步触发输出功能

 负载保护		 适应强电感性负载

 ATS模式防止电流和电压冲击

 输出端保护防止上电冲击和静电损伤

 输出端状态可控，支持开机更换负载

 交互能力		 白色背光FSTN LCD显示器

 必备的LED状态指示器

 与显示器和指示器配合的34键交互型功能键盘操作

 标准规格		 提供用户手册和设备驱动程序

 标准半宽2U台式仪器，交流220V/50Hz供电

 散热方式：强制风冷

 外形尺寸WHD=21788317mm

标准配置：	 仪器		F2002				程控精密电流源						1台

 附件		F20001				二线式输出转接器					1台

CAB-AC250-052 		3芯交流250V/5A GB1002电源线2m	1条

CAB-232-DCE-3		RS-232C标准直通电缆3m				1条

CAB-USB-AB-2		USB标准A-B电缆1.8m				1条

FUSE5G-25050		250V/0.5A玻封保险管，520mm		2只

 光盘		CD-F2002	 		产品信息光盘						1片

内含：				MAN-F2002		用户手册电子文档

DRV-USB232		USB接口驱动程序

Single Page Introduction – F2002

Document Number：	SPI-F2002-REV01-2012

Model:	F2002

Name:	Programmable Precision Current Source

Key words:	105mA/105V/11W, strong inductance load, configurable compliance

high precison, high stability, high resolution, programmable

Warranty:		Limited warranty for a period of 1 year from delivery

Calibration:	Free for the first time calibration

Measuration: 	Without force measuration product, payment for Certification of China

Standard:		Follows GB-1002, EIA-232-E, USB 1.1 Standard and IEC 825-1

Specification:	 Output Current		 Range of 0 to 105mA DC, quasi-bipolar

 1A step resolution, (1/105, 000 F.S.)

 Accuracy of  (0.015% settings+1A)

 Linearity of better than 0.005%

 Temperature coefficient of better than 3ppm/C

 Short term stability of better than 10ppm/h

 Noise and ripple of 1A RMS typical (Low Noise mode)

 DC output resistance >500M

 Floating output at I, III region, with internal relay switching

 Compliance		 Range of 0 to 105V DC

 0.1V step

 Accuracy of  (0.5% Settings+0.5V)

 Output Power		 Maximum output power of 11W

 Automation		 RS-232C interface as standard configuration

 Internal USB-RS232 converter

 Synchronous triggering output with optional delay time

 Load Protection	 Capability of driving strong inductance load

 ATS mode to avoid current or voltage spikes

 Output protection to avoid power-on or static shocks

 Controllable output status allows power-on load connections

 Human Interface	 FSTN LCD display with white back lighting

 LED indicators

 34 functional keyboard operation

 Standard Spec.		 User’s manual and device drivers

 Standard half rack 2U instrument, AC 220V/50Hz supply

 Forced air cooling

 Dimension of WHD=21788317 mm

Standard Configuration:

 Instruments		F2002 				Programmable Precision CS		1

 Accessories		F20001				2-Wire Output Converter			1

CAB-AC250-052		GB1002 AC 250V/5A 3-conductor	1

power cord of 2m long

CAB-232-DCE-3		Std. RS-232C straight 	through		1

cable of 3m long

CAB-USB-AB-2		Std. USB A-B cable of 1.8m long	1

FUSE5G-25005		250V/0.5A glass fuse, 520mm	2

 CDs				CD-F2002			Product information CD			1

which contains:		MAN-F2002 user’s manual e-documents

DRV-USB232 interface driver